Acoustic Transparency Comparisons of Hearing Aid Retention Accessories

| RESP 50 - volume control taped so it remains stable throughout evaluation |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Set-up of hearing aid with retention accessory | 200 | 250 | 315 | 400 | 500 | 630 | 800 | 1000 |
| Hearing aid only | 63 | 65 | 71 | 72 | 75 | 79 | 84 | 90 |
| Hearing aid on false ear (pinna effect) | 64 | 68 | 70 | 73 | 75 | 78 | 82 | 90 |
| Hearing aid on false ear with Ear Gear | 61 | 67 | 70 | 72 | 76 | 79 | 83 | 89 |
| Hearing aid on false ear with Hanna Andersson cap | 63 | 67 | 70 | 73 | 75 | 78 | 82 | 91 |
| Hearing aid on false ear with Silkawear cap | 63 | 67 | 70 | 73 | 75 | 78 | 82 | 91 |
| Hearing aid on false ear with Hearing Henry headband | 63 | 67 | 70 | 69 | 74 | 79 | 82 | 86 |

| Set-up of hearing aid with retention accessory | 1250 | 1600 | 2000 | 2500 | 3150 | 4000 | 5000 | 6300 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Hearing aid only | 99 | 93 | 87 | 87 | 86 | 83 | 74 | 55 |
| Hearing aid on false ear (pinna effect) | 99 | 94 | 91 | 95 | 91 | 85 | 75 | 56 |
| Hearing aid on false ear with Ear Gear | 98 | 91 | 88 | 93 | 89 | 83 | 73 | 54 |
| Hearing aid on false ear with Hanna Andersson cap | 100 | 96 | 93 | 95 | 93 | 88 | 79 | 56 |
| Hearing aid on false ear with Silkawear cap | 100 | 96 | 93 | 95 | 93 | 88 | 79 | 56 |
| Hearing aid on false ear with Hearing Henry headband | 91 | 84 | 79 | 84 | 88 | 79 | 68 | 48 |

| GAIN IN dB |  |  |  |  |
| --- | --- | --- | --- | --- |
| Set-up of hearing aid with retention accessory | Average OSPL90 | Average gain@50 | Average gain@60 | Frequency Range |
| Hearing aid only | 109 | 30 | 30 | 300-5600 |
| Hearing aid on false ear (pinna effect) | 109 | 32 | 33 | 375-5600 |
| Hearing aid on false ear with Ear Gear | 109 | 31 | 31 | 315-5340 |
| Hearing aid on false ear with Hanna Andersson cap | 109 | 34 | 34 | 420-5600 |
| Hearing aid on false ear with Silkawear cap | 109 | 34 | 34 | 420-5600 |
| Hearing aid on false ear with Hearing Henry headband | 108 | 24 | 24 | 200-5340 |

| HARMONIC DISTORTION |  |  |  |
| --- | --- | --- | --- |
| Set-up of hearing aid with retention accessory | 500 Hz | 800 Hz | 1600 Hz |
| Hearing aid only | 2 | 1 | 2 |
| Hearing aid on false ear (pinna effect) | 3 | 2 | 2 |
| Hearing aid on false ear with Ear Gear | 3 | 1 | 1 |
| Hearing aid on false ear with Hanna Andersson cap | 2 | 1 | 1 |
| Hearing aid on false ear with Silkawear cap | 2 | 1 | 1 |
| Hearing aid on false ear with Hearing Henry headband | 2 | 1 | 2 |

Phonak Piconet2 (P2AZ) hearing aid used for assessment. Set volume backed off slightly from full on and taped into position.

Refer to next page for photos of acoustic transparency assessment set up.

Performed October 2013 by Vickie Anderson, AuD and Karen Anderson, PhD using an AudioScan electroacoustic hearing aid analyzer.

Acoustic Transparency Assessment of Hearing Aid Retention Accessories

Hearing aid on false ear (to account for pinna effect) with Ear Gear

Hearing aid covered by Silkawear cap and Hanna Andersson pilot cap

Hearing aid covered by Hearing Henry Headband